## Sheet1
Content Approval: Date: Brian Ross See M-Files Authorizer: Date: Ann Holt See M-Files Distributed: M-Files |  |  |  | References: 21 CFR 820.200 Servicing 21 CFR 803 Medical Device Reporting
Safe and effective use of the EOGas4 Ethylene Oxide Gas Sterilizer (Sterilizer) begins with a thorough understanding of all sections of the Owner’s/User’s Manual (Manual) that is provided with the Sterilizer. It is strongly suggested to read the entire Manual prior to performance of Inspection, Service, and Maintenance, and using the Sterilizer. Proper preventative maintenance according to the Manual cannot be over emphasized.
The following specific features pertaining to Preventative Maintenance instructions are discussed in the Manual and are itemized here. Response to these questions and statements should be recorded and retained along with other important Sterilizer related documents in a Sterilizer Equipment File for reference and review should the need arise.
Sterilizer & Customer Information:
Date of Service:
Service Type:
Performed By (Name of Authorized Technician):
ACT Record; Sterilizer Model Number: |  | AN4000.01
ACT Record; Sterilizer Lot Number:
ACT Record; Customer Number:
ACT Record; Customer Name:
ACT Record; Customer Location:
ACT Record; Customer Contact Name:
ACT Record; Customer Contact Phone:
Equipment Used: | MAKE | MODEL | SERIAL NUMBER |  | CALIBRATION DUE DATE
FLOW METER | King | 75301112C17 | 83790121002 |  | 2023-01-22
VACUUM GAGE | Reed | R3030 | 9147573 |  | 2023-01-16
RELATIVE HUMIDITY GAGE | Traceable | ZM-90080-06 | 210867166 |  | 2023-01-22
TEMPERATURE GAGE | Fluke | 325 | 57312124MV |  | 2023-08-03
Exterior Exhaust Location Recommendations: |  |  | ACCEPTED | COMMENTS
1. Confirm that the exhaust outlet does not terminate near a common area or within 25’ (7.6m) of any building intake source. (ANSI/AAMI ST41 Guidance)
2. Confirm that the sterilizer is connected to a dedicated exhaust run (Note any changes or updates to the exhaust system)
3. Locate & Confirm exhaust outlet is free of visible clogs or obstructions, and located such to minimize rain or storm water incursion. Document location by Photograph or Sketch.
Customer & Equipment Record Verification: |  |  | ACCEPTED | COMMENTS
1. Verify Customer & Sterilizer Records above. (Note any changes or updates) |  |  | Y
As Found Sterilizer Observations: |  |  | ACCEPTED | COMMENTS
1. Document Sterilizer Installed Location by means of photograph or Sketch. |  |  | Y
2. Document Sterilizer Exhaust hose configuration by means of Photograph or Sketch. |  |  | Y
3. Verify cabinet is level and stable. |  |  | Y
4. Verify that the sterilizer is connected to a dedicated, unswitched, grounded power receptacle of the correct rating with no GFCI and that the Smart Power, power conditioner is in use. Document by Photo |  |  | NA | To be checked at next service visit
5. With door closed; Verify door lock engagement & about 2mm door sensor gap. |  |  | Y
6. Inspect cabinet door hinges for wear. |  |  | Y
7. Verify Door Open Alarm Function (Proximity sensor position). |  |  | Y
8.Using a polarity tester verify that when energized, the door magnet and solenoid are oriented so that the magnetic poles are aligned. (If the magnet is North pole up, door solenoid should be North pole up) |  |  | Y
9. With door closed; Inspect door seal for visible signs of gaps or leaks. |  |  | Y
10. Suspend calibrated temperature sensor in center of cabinet. Route sensor to outside of cabinet ensuring minimal seal deformation upon door close. |  |  | Y
11. Turn On Sterilizer, press start, record Pump Hours. (Photograph screen if possible)
12. Allow Sterilizer to heat to standard operating temperature. (50°C) |  |  | Y
13. Perform Leak Check on Exhaust Hose as directed by Manual. |  |  | Y
14. Inspect Circulation Fan for rpm, bearing bind, or noise during operation. (Recommended replacement interval is 3 yrs.) |  |  | Y
15. Install test bag onto Purge Probe, allow Sterilizer to Purge; Photograph fully purged Sterilization Chamber. |  |  | Y
16. Confirm Bag Label Printer Function (Document sample Label by photograph) |  |  | Y
17. Run Sterilizer Diagnostics; Verify Purge V, Vent V, & Back P all indicate "0.0" at start of sequence. (Photograph screen) |  |  | Y
As Found Sterilizer Observations (continued): |  |  | MEASURED | ACCEPTED | COMMENTS
18. Record Relative Humidity in Area of Sterilizer. (% RH) |  |  |  | Y | IF LESS THAN 35%RH, USE OF HUMIDICHIP IS RECOMMENDED
19. Record Temperature in Area of Sterilizer. (°C) |  |  |  | Y | IF LESS THAN 20°C, STERILIZER MAY BE SLOW TO REACH SET TEMPERATURE AFTER LOADING
20. Cabinet temperature sensor. Acceptance Criteria: 50°C ± 0.5°C |  | NIST Sensor: |  | N | CALIBRATION REQUIRED
Sterilizer:
21. With Sterilizer operating in Diagnostics Mode; Record sensor values (Photograph screen) | Acceptance Criteria: | Purge V (> 5.0) |  | N
|  | Vent V (> 0.25) |  | N
|  | Back P (< 15) |  | Y
|  | Fan (> 70%) |  | N
22. With Sterilizer operating in Diagnostics Mode; Record Purge Probe Valve Vacuum Value | Acceptance Criteria: | < -0.22 psi |  | N
23. Connect Flow Meter to Purge Probe Tube; Record flow rate in Diagnostics Mode. | Optimal Purge Pump Flow Performance Criteria: | > 1.34 cfm |  | Corrective Action Recommended | MINIMIUM AIR FLOW REQUIRED FOR 10 CABINET AIR CHANGES PER HOUR IS 0.84 CUBIC FEET PER MINUTE (cfm)
24. Connect Flow Meter to Vent inlet connector at rear upper left corner of cabinet interior and record flow rate in Diagnostics Mode. | Optimal Vent Pump Flow Performance Criteria: | > 1.34 cfm |  | Corrective Action Recommended | MINIMIUM AIR FLOW REQUIRED FOR 10 CABINET AIR CHANGES PER HOUR IS 0.84 CUBIC FEET PER MINUTE (cfm)
As Found Sterilizer Observations (continued): |  |  | MEASURED | ACCEPTED | COMMENTS
25. Disconnect Power Cord From Sterilizer. |  |  |  | Y
26. Remove Top Cover from Sterilizer. (Note Any Observations & Photograph Interior, ensuring that Lot Label is captured within image) |  |  |  | Y
27. Connect Vacuum Meter to Inlet tube of Vent Pump. |  |  |  | Y
28. Ensure that no unauthorized personnel are within 3 feet of the Sterilizer; Connect Power Cord to Sterilizer. |  |  |  | Y
29. Turn on Power to the Sterilizer; Press Start to Initialize Self Test; record Vent Pump Obtainable Vacuum. | Acceptance Criteria: | <-5.42 psi |  | N
30. Disconnect Power Cord From Sterilizer. |  |  |  | Y
31. Reconnect Vent Pump Tube Harness. |  |  |  | Y
32. Connect Vacuum Meter to Inlet tube of Purge Pump. |  |  |  | Y
33. Ensure that no unauthorized personnel are within 3 feet of the Sterilizer; Connect Power Cord to Sterilizer. |  |  |  | Y
34. Turn on Power to the Sterilizer; Press Start to Initialize Self Test; record Purge Pump Obtainable Vacuum. | Acceptance Criteria: | <-5.42 psi |  | N
35. Disconnect Power Cord From Sterilizer. |  |  |  | Y
List Sterilizer Repair Materials & Activities (if repair is required):
Observed Issue:
Billable Repair Materials Installed
Material Description | Quantity | Part Number | Lot Number
Description of Repair (Root Cause Identification):
Preventative Maintenance Activities |  |  | ACCEPTED | COMMENTS
1. Clean Sterilizer inside & outside using cotton cloth and mild soap and water solution. |  |  | Y
2. Replace head unit cooling fan & Head filters. |  |  | Y
3. Based upon Pump test data in "As Found" table above; Replace any pumps which do not meet performance requirements. Note: Proactive replacement is recommended at limit +5%. |  |  | Y
4. Based upon Purge Probe Valve test data in "As Found" table above; Replace any Valves which do not meet performance requirements. Note: Proactive replacement is recommended at limit +5%. |  |  | Y
5. Inspect Internal Tubing for signs of deterioration, discoloration, hardening; Replace if necessary. |  |  | Y
6. Inspect Sterilizer Internal Tubing for presence of hose clamps at all exhaust tube barb fitting locations. |  |  | Y
7. Inspect tube and cable routing within sterilizer top assembly; ensure that tubes and cables cannot rub adjacent objects due to vibration. Tie tubes and cables as needed. |  |  | Y
8. Inspect Pump/Cage clearance; ensure that minimum 2mm clearance is present surrounding pumps. |  |  | NA | No pump cages on EO4
9. Confirm exhaust bulkhead fitting lock ring is tight. |  |  | Y
10. Check Internal Tubing Fittings for Leaks. |  |  | Y
11. Perform Leak Check on Exhaust Hose as directed by Manual. |  |  | Y
12. Confirm Door Open Alarm function. |  |  | Y
13. Replace Bag Label Printer roll. |  |  | Y
No-Charge Preventative Maintenance Materials Used:
Material Description | Quantity | Part Number | Lot Number
Cabinet Filter | 3 | SN7471
Bag Label Printer Roll | 1 | AN1000.97
Internal Tubing Kit |  | SN55000001
As Left Sterilizer Observations: |  |  | MEASURED | ACCEPTED | COMMENTS
1. Ensure that no unauthorized personnel are within 3 feet of the Sterilizer; Connect Power Cord to Sterilizer. |  |  |  | Y
2. Turn on Power to the Sterilizer; Press Start to Initialize Self Test. |  |  |  | Y
3. Connect Flow Meter to Purge Probe Tube; Record flow rate in Diagnostics Mode. | Optimal Purge Pump Flow Performance Criteria: | > 1.34 cfm |  | Corrective Action Recommended | MINIMIUM AIR FLOW REQUIRED FOR 10 CABINET AIR CHANGES PER HOUR IS 0.84 CUBIC FEET PER MINUTE (cfm)
4. Connect Flow Meter to Vent inlet connector at rear upper left corner of cabinet interior and record flow rate in Diagnostics Mode. | Optimal Vent Pump Flow Performance Criteria: | > 1.34 cfm |  | Corrective Action Recommended | MINIMIUM AIR FLOW REQUIRED FOR 10 CABINET AIR CHANGES PER HOUR IS 0.84 CUBIC FEET PER MINUTE (cfm)
5. Disconnect Power Cord From Sterilizer. |  |  |  | Y
Complete applicable Steps 6 through 13 if Purge or Vent Pumps have been replaced during service. |  |  |  | N/A
6. Connect Vacuum Meter to Inlet tube of Vent Pump. |  |  |  | N/A
7. Ensure that no unauthorized personnel are within 3 feet of the Sterilizer; Connect Power Cord to Sterilizer. |  |  |  | N/A
8. Turn on Power to the Sterilizer; Press Start to Initialize Self Test; record Vent Pump Obtainable Vacuum. | Acceptance Criteria: | <-5.42 psi | N/A | N/A
9. Disconnect Power Cord From Sterilizer. |  |  |  | N/A
10. Connect Vacuum Meter to Inlet tube of Purge Pump. |  |  |  | N/A
11. Ensure that no unauthorized personnel are within 3 feet of the Sterilizer; Connect Power Cord to Sterilizer. |  |  |  | N/A
12. Turn on Power to the Sterilizer; Press Start to Initialize Self Test; record Vent Pump Obtainable Vacuum. | Acceptance Criteria: | <-5.42 psi | N/A | N/A
13. Disconnect Power Cord From Sterilizer. |  |  |  | N/A
Complete Steps 14 & 15 if purge probe has been replaced. |  |  |  | N/A
14. Ensure that no unauthorized personnel are within 3 feet of the Sterilizer; Connect Power Cord to Sterilizer. |  |  |  | Y
15. With Sterilizer operating in Diagnostics Mode; Record Purge Probe Valve Vacuum Value | Acceptance Criteria: | < -0.22 psi | N/A | N/A
16. Turn On Sterilizer, press start, record Pump Hours. (Photograph screen) |  |  |  | 17941
17. Run Sterilizer Diagnostics; Verify Purge V, Vent V, & Back P all indicate "0.0" at start of sequence. (Photograph screen) |  |  |  | Y
18. Allow Diagnostics to continue; Record sensor values (Photograph screen) | Acceptance Criteria: | Purge V (> 5.0) |  | N
|  | Vent V (> 0.25) |  | N
|  | Back P (< 15) |  | Y
|  | Fan (> 70%) |  | N
19. Disconnect power cable from sterilizer. |  |  |  | Y
20. Confirm Vent & Exhaust Tubes are secured to tube support post. |  |  |  | Y
21. Confirm minimum 2mm pump/cage clearance. |  |  |  | NA | No pump cages on EO4
22. Confirm no tools or debris from service remain within sterilizer head assembly. |  |  |  | Y
As Left Sterilizer Observations (Continued): |  |  | MEASURED | ACCEPTED | COMMENTS
23. Document Sterilizer head assembly by means of Photograph (with Lot #). |  |  |  | Y
24. Inspect tubes and cables for potential rub locations. |  |  |  | Y
25. Confirm presence of mechanical hose clamps at all exhaust tube fittings (9 locations) |  |  |  | Y
26. Install Top Cover. |  |  |  | Y
27. Run Sterilizer Diagnostics; Verify Purge V, Vent V, & Back P all indicate "0.0" at start of sequence. (Photograph screen) |  |  |  | Y
28. Allow Diagnostics to continue; Record sensor values (photograph screen if possible) | Acceptance Criteria: | Purge V (> 5.0) |  | N
|  | Vent V (> 0.25) |  | #REF!
|  | Back P (< 15) |  | Y
|  | Fan (> 70%) |  | N
29. Cabinet temperature sensor Calibration Verification. Acceptance Criteria: 50°C ± 0.5°C |  | NIST Sensor: |  | Y
Sterilizer:
30. Verify Sterilizer runs quietly & without rattles. |  |  |  | Y
31. Capture Photograph of fully purged Sterilization Bag. |  |  |  | Y
32. Verify successful start of 3 Hour Cycle. |  |  |  | Y
33. Verify successful start of 5 Hour Cycle. |  |  |  | Y
34. Verify successful start of 12 Hour Cycle. |  |  |  | Y
35. Verify successful start of 24 Hour cycle. |  |  |  | Y
36. Confirm Function of Label Printer (Photograph Printed Label) |  |  |  | Y
37. Attach Next Service Due Label to cabinet door. |  |  |  | Y
38. Document Sterilizer "As Left" Location by means of photograph or Sketch. |  |  |  | Y
Service Finalization Activities: |  |  | Yes | No | COMMENTS
1. Discuss importance of regular Preventive Maintenance with Sterilizer Owner. |  |  | Yes
2. Record all available Sterilizer Operator's Names and Key Op Certification Status. |  |  |  | No | Contact Andersen rep for list of personnel
Operator Name: |  | Key Op Certified:
Operator Name: |  | Key Op Certified:
Operator Name: |  | Key Op Certified:
Operator Name: |  | Key Op Certified:
Operator Name: |  | Key Op Certified:
Operator Name: |  | Key Op Certified:
If the above steps have been completed successfully, the Sterilizer can now be considered ready for use, or if required, ready for qualification testing. |  |  | Customer Signature:
Technician Signature:
Final Approval - To Be Filled In by Andersen Products Reviewing Manager
Final Review Activity: |  |  | Yes | No | COMMENTS
All fields of Report properly populated
Technician & Customer Acceptance Signatures present
Human Use Sterilizer
Observed or Customer communicated Issue; Review for Determination of Reportable Incident
CFF Required
CFF Created, (If Yes, Enter CFF Number in Comments Field)
APPROVED: Liz Krewson SIGNED: See M files |  |  |  | DATE: | See M Files
Revision History
| ECN # |  | Date | Rev. | Description of Revision
| 170008-2017 |  | 29-12-17 | 0 | Initial Release
| 180001-2018 |  | 2018-08-10 | 1 | Update for ISO 13485:2016
| 2020038-2020 |  | See M-Files | 2 | Updated for SCAR 2020105 submitted by manufacturer, Item # 4 added
AS FOUND PHOTOGRAPHS:
AS LEFT PHOTOGRAPHS: